## Sheet1
| MODBUS RS485 RTU no parity 8 bit data 2 stop bits | SONOWELD - aluminum case generators
register | parameter | dim | bit 0 | bit 1 | bit 2 | bit 3 | bit 4 | bit 5 | bit 6 | bit 7 | bit 8 | bit 9 | bit 10 | bit 11 | bit 12 | bit 13 | bit 14 | bit 15 | bit 16
0 | BEAT counter | cnt
1 | Start frequency | accuracy 10 Hz
2 | Output power set point | W
3 | Span set point | ru
4 | Spare | ru
5 | Regulator Ki | ru
6 | Regulator Kv | ru
7 | I max | mA
8 | Modbus address | adr
9 | Command word | bit field | bit 0: RESET; bit 1: START; bit 2: STOP; bit 3: SCAN; bit 4: WRITE; bit 5: TEST; bit 6: Spare; bit 7: Spare
10 | Status | ASCII code | bit 0: 0x30 stop; bit 1: 0x31 start; bit 2: 0x32 overcurrent; bit 3: 0x33 overvoltage; bit 4: 0x34 overheat; bit 5: 0x35 Load error; bit 6: 0x36 External overheat; bit 7: 0x37 Frequency error
11 | Analog set point | W
12 | Output Current | mA
13 | Output voltage | V
14 | Output PWM | %x10
15 | Output power | W
16 | Spare | ru
17 | Frequency actual | accuracy 10 Hz
18 | Frequency shift | accuracy 10 Hz
19 | Spare | ru
20 | Spare | ru
21 | Spare | ru
22 | SP Phase | ru
23 | Modbus peed | code | bit 0: 0-19200b; bit 1: 1-38400b; bit 2: 2-115200b
24 | Nominal power | W
25 | SP amplitude | V
26 | Current scale | ru
27 | External temperature | °C x 10
28 | Nominal amplitude | V
29 | Relay status | bit field | bit 0: Relay HF/RF; bit 1: Relay Error
30 | Spare | ru
31 | Phase | ru
32 | Peak power | W
33 | Peak power plus | W
34 | Peak power minus | W
35 | Energy | J
36 | Energy plus | J
37 | Energy minus | J
38 | Time | ms x 10
39 | Time plus | ms x 10
40 | Time minus | ms x 10
41 | Welding status | ASCII code | bit 0: 0x30 OK; bit 1: 0x31 limit peak power; bit 2: 0x32 limit energy; bit 3: 0x31 limit time
42 | Last peak power | W
43 | Last energy | J
44 | Last time | ms x 10
45 | Delay | ms x 10
46 | Hold | ms x 10
47 | After burst | ms x 10
48 | Firmware version | ru
49 | Button mode | code | bit 0: 0-off; bit 1: 1-push; bit 2: 2-latch; bit 3: 3-triger
50 | Output power filtered | W
51 | External temperature threshold | °C x 10
………
100 | Scan
292 | data
……….